## Slide 1: Hands on data: Chemometrics
Mathematical modelling
27th October, Mercedes Bertotto

## Slide 2: Workshop agenda: Chemometrics
Chemometrics: Definition (5 min)
General workflow for chemometrics (2 min)
Exploratory analysis and Outlier Detection (5 min)
Pretreatments on raw spectra (5 min)
Feature selection (5 min)
Cross validation or Data Split (1 min)
Calibration and validation of classification models (7 min)
Hands on Data (30 min)

## Slide 3: Part 1: Chemometrics: Definition
"Delve deep enough into anything, and you will find mathematics" - Dean Schlicter

## Slide 4: Chemometrics
“Chemometrics is the chemical discipline that uses mathematical, statistical, and other methods employing formal logic to design or select optimal measurement procedures and experiments, and to provide maximum relevant chemical information by analyzing chemical data“. D. L. Massart (1941-2005)
Calibration
Pattern recognition
Design of Experiments
Notes: "Métodos multivariados de analysis de datos aplicados a datos de interés químico.“
Otro decia que chemometrics Is what chemometricians do.

## Slide 5: Statistics versus chemometrics
Focus on distributions
Inference based on hypothesis tests of parameters
Density estimation
Mandatory courses at every biological education
Focus on individual samples
Inferences based on future performance
Clustering
Voluntary course at only few institutions
Notes: Chemometrics is the science of extracting information from chemical systems by data-driven means. Chemometrics is inherently interdisciplinary, using methods frequently employed in core data-analytic disciplines such as multivariate statistics, applied mathematics, and computer science, in order to address problems in chemistry, biochemistry, medicine, biology and chemical engineering. In this way, it mirrors other interdisciplinary fields, such as psychometrics and econometrics.

## Slide 6: Part 1: Chemometrics
NIR spectra are of high dimension
The data is highly correlated and obscured by the presence of overlapping absorbances, harmonics, and combination bands
Spectra are often complicated by light scattering and other physical effects
Multivariate methods (chemometrics) are required to address these issues
“Linear algebra is the language of chemometrics. To understand most chemometric techniques, a basic understanding of linear algebra is required." (Wise and Gallagher, 1998)

## Slide 7: Part 2: General workflow for chemometrics

## Slide 8: GOOD MODELING PRACTICES (GMP)
Collection of a set of representative samples
Measurement of spectra of the entire collection
Creating a calibration set
Creating a validation set
Analysis of all samples by a reference method
Establishment of a predictive model through the application of multidimensional statistical methods
Validation of the model on the validation set
Add new samples to the collection
Were the control samples analyzed accurately?
Does the equipment work correctly?
Instrument repair/check
Are the control samples representative?
Check the reference dosage
Serial analysis of unknown samples
YES
YES
YES
Notes: "BPMs (Best Practices in Chemometrics) aim to explain and describe the stages to be carried out when implementing a general analytical model. The stages presented here are based on the existence of predictive variables in the form of spectra (signals of various origins: IR spectroscopy, Raman, fluorescence, etc.).
Commencement: Study the objective of the work and the best way to achieve it. The data collection stage must also adhere to certain rules to be effective.
From the collected data, a predictive model will be constructed while taking precautions to ensure practical validity. The established model can be applied in practice, and its validity must be regularly verified.
Pre-data collection stage: Reflect on the analysis strategy to be implemented to solve a problem. It should be as simple as possible: there are industrial situations where it is recommended to directly test the product's compliance with a reference spectral collection. Spectral distance.
Study relevance, accuracy, the domain of application, etc. Choose the reference method. Select a sufficient number of representative samples from the population. These samples will then be divided into two sets: calibration and validation. A third set of data can be extracted from the initial samples or acquired specifically for testing purposes.
The results of the analysis of the calibration samples are used to create a regression or discrimination model. This model is then applied to the results of the validation samples and tested with the test set. After validation, the predictive model can be used in routine.
The proper functioning of the spectroscope and the chemometric model should be regularly monitored and potentially adjusted."

## Slide 9: Data Set Preparation
GOOD MODELING PRACTICES (GMP)
Notes: Samples: they must be measured under standardized conditions. They should be representative of what will be measured later.
Variability can be contained in many factors, such as biological sample variability, geographical origin, flowering date, or harvest date.
The first stage in establishing a predictive model involves splitting the spectral collection into two sets: learning and test sets. The first will be used to fine-tune the predictive model, and the test set will be used to estimate the quality of the predictive model.
The term 'set' refers to spectra selected using an alternative spectroscopic method, and the term 'value' refers to the reference method.
The division can be performed using different algorithms that we will see later on.

## Slide 10: Part 3: Exploratory analysis and Outlier Detection

## Slide 11: Reducing dimensionality
Identifying and Removing Irrelevant Variables to:
Enhance Computational Performance.
Improve Model Interpretability and Result Comprehension.
Avoid overfitting and colinearity.
One way to solve the problem: Transforming the initial set of variables into a lower-dimensional set while retaining most of the information
Notes: More representative observations improve the model, but a higher number of variables not necessarily
To solve the problem: Transforming the initial set of variables into a lower-dimensional set while retaining most of the information

## Slide 12
Objectives:
Exploration/ Description
Dimensionality Reduction
Preparation and Cleaning (Outlier Identification, Noise)
Discrimination of Individual Groups
Determination of Relationships Among Individuals
Quantification of Variable Correlations (Individual-Variable Duality)
Preliminary Stage for Further Chemometric Treatment
Principal Component Analysis (PCA)
Notes: Un ACP sirve en primer lugar, para explorar o para describir un set de datos multivariados que presenta una dimensión elevada, es decir un gran número de variables. (Reducción de la dimensionalidad). La visualización de los resultados en un espacio reducido, un plano por ejemplo, permite sintetizar la info contenidoa en un juego de datos.
Tambien sirvE para limpiar y preparar el set de datos antes de efectuar otro tratamiento. La identificación de outliers, se realiza de modo multivariado. El ACP constituye una etapa prealable, indispensable para todo analisis de datos multivariados. Extraer info de un set de datos, separando el ruido tanto aleatorio como sistemático.
El ACP también es una herramienta de identificación de grupos de individuos similares dentro de un juego de datos. La representación de la proximidad entre individuos muestra a los individuos que se comportan de forma similar en relación a las variables que fueron medidas.
De igual manera que la proximidad entre los individuos es puesta en evidencia, la dualidad individuos-variables de un ACP permite cuatificar las correlaciones entre las variables medidas de un juego de datos.
Dans certains cas, l’ACP constitue une étape préalable à un traitement
chimiométrique ultérieur.
Par exemple, une ACP peut être réalisée avant une classification non supervisée,
de façon à réduire la dimensionnalité initiale du jeu de données.
Pour certaines méthodes de discrimination comme SIMCA (Soft Independent
Modeling of Class Analogies), la modélisation de chaque classe se réalise à l’aide
d’une ACP [3].
Enfin, certaines applications de supervision de procédés utilisent le suivi dans le
temps des scores issus d’une ACP pour la détection des dérives ou d’événements
anormaux survenant au cours du déroulement du procédé.

## Slide 13
Principal Component Analysis (PCA)
Notes: "Describe the variation in multivariate data through a small number of uncorrelated variables (Principal Components, PC). Combines the original variables into a set of new variables in a way that the most information in the original data prevails in the new variables. Reduces dimensionality while preserving the maximum variance. The PCs are uncorrelated and are ordered in a way that the first ones retain most of the variation present in the original set of variables.
New Latent Variables are constructed, variables that had not been measured
The original variables are proyected into new Latent Variables

## Slide 14
Error de proyección total: 11.3
Error de proyección total: 34,5
Principal Component Analysis (PCA)
Notes: Si usamos el umbral de varianza para eliminar Y, el número de características, también conocido como la dimensionalidad del conjunto de características, se reduce en uno. Geométricamente, esto corresponde a los datos bidimensionales x-y que se proyectan en el eje de la característica restante X.. Una medida de la información que se pierde es el error de proyección, que es la distancia total entre los puntos de datos originales y proyectados. Para este conjunto de características, está claro que el error incurrido al proyectar en el eje x es mucho más pequeño que si en su lugar hubiéramos proyectado en el eje y. Al observar una muestra del segundo conjunto de características, está claro que los errores de proyección serán grandes tanto para el eje x como para el eje y, y ninguno de los ejes minimizará el error.
When using a variance threshold to eliminate a feature Y, the number of features, also known as the dimensionality of the feature set, is reduced by one. Geometrically, this corresponds to two-dimensional data, x-y, being projected onto the axis of the remaining feature X. A measure of the lost information is the projection error, which is the total distance between the original data points and the projected ones. For this feature set, it's clear that the error incurred when projecting onto the x-axis is much smaller than if we had projected onto the y-axis instead. When looking at a sample from the second feature set, it's evident that the projection errors will be large for both the x and y axes, and neither axis will minimize the error."

## Slide 15
Principal Component Analysis (PCA)
Notes: "However, you can minimize the projection error by defining a new axis. The direction of this special axis is called the first principal component of the feature set. Principal components have a couple of properties that make them especially useful for dimensionality reduction.
They determine the axis direction that minimizes the projection error, and as a result, this is also the direction of maximum variance. In this example, since the principal component axes capture most of the variance, you can use it to define a new feature, P. You can then replace X and Y with P, reducing the dimensionality of the feature set by one. The values of the new entity are determined by projecting the values of the original entity onto the new axis and measuring the signed distance from the origin to each projection. These values are commonly known as the scores of the principal components."

## Slide 16
Principal Component Analysis (PCA)
Notes: So far, we've only discussed the first principal component. The total number of principal components is equal to the dimensionality of the feature set. The second principal component is the direction of maximum variance that the first principal component doesn't account for; therefore, its direction is perpendicular or orthogonal to the first principal component. In more than two dimensions, the principal components are determined in decreasing variance order as follows. Start by finding the first principal component. The second principal component lies in the direction of maximum variance among all directions orthogonal to the first principal component. The third principal component is the direction of maximum variance among all directions orthogonal to the first two principal components. This process continues until only one orthogonal direction remains, which becomes the final principal component.
The result is a new coordinate system where the original axes have been rotated to maximize variance along the new axes.

## Slide 17: How PCA works
Notes: In this example, the variance of variable A is 0.67, and the variance of the second variable is 0.25. The covariance between the two variables is 0.55, as reflected in the main diagonal of the matrix.
Because they are square and symmetric, covariance matrices are diagonalizable, meaning that an eigenvalue decomposition can be performed on the matrix. This is where PCA finds the eigenvectors and eigenvalues for the dataset.
An eigenvector of a linear transformation is a vector (other than zero) that changes by a scalar multiple of itself when the related linear transformation is applied. The eigenvalue is the scalar associated with the eigenvector.
Eigenvalues are important because they provide a ranking criterion for the newly derived variables (axes). The principal components (eigenvectors) are sorted in descending eigenvalue order. The principal components with the highest eigenvalues are "selected first" as principal components because they represent the most variation in the data.
You can specify to return nearly as many principal components as there are variables in your original dataset (usually up to n-1, where n is the number of original input variables), but most of the variance will be accounted for in the top principal components.
After identifying the principal components of a dataset, the observations from the original dataset need to be transformed into the selected principal components.
To convert the original points, we create a projection matrix. This projection matrix is simply the selected eigenvectors concatenated into a matrix. We can then multiply the matrix of our original observations and variables by our projection matrix. The result of this process is a transformed dataset, projected into our new data space, composed of our principal components!

## Slide 18: PCA: Matrix decomposition
(N x P)
(N x A)
(P x A)
Notes: The initial data table, i.e., the (N x P) dimension matrix X, is expressed as the product of the matrix T of scores (or factorial coordinates), with dimensions (N x A), where A is the number of retained principal components, and the transposition of the matrix P of loadings, with dimensions (P x A). Additionally, the matrix E, representing the residual variance, with dimensions (N x P), is included. It's worth noting that A, the number of retained principal components for the model, is typically much smaller than N and P. In fact, it's a reduction in dimensionality.
Scores: Each Principal Component represents a specific percentage of variance explained by the model, in relation to the total variance in the dataset, which sums up to 100%. The Principal Components are always ordered in decreasing order of the explained variance percentage.
Loadings: They represent the initial variables in the vector space constructed around the Principal Components. The distance of the projection of an initial variable onto the Principal Components corresponds to a loading or factorial contribution, meaning its weight or contribution or influence in building the PCs. The cosine of the angle formed between an initial variable and a PC corresponds to the correlation between both vectors.
Remember that a Principal Component is a linear combination of initial variables. The loadings correspond to the parameters 'p' in this equation.

## Slide 19: PCA: Matrix decomposition
Notes: The breakdown of PCA is also provided for the specific case of spectral data. A measured spectrum belonging to the spectral base is decomposed component by component as the product of a score (or factorial coordinate) and a loading, which is seen as a spectrum. This loading represents the useful information. The residual variance matrix contains the unexplained part of the signal, i.e., the noise.
Remember that the optimization criterion for PCA is to maximize the variance of the X matrix transformed by the Principal Components (PCs) (2). This is equivalent to diagonalizing the dot product X'X, which corresponds to the covariance matrix (3).
From this algebraic operation, the calculation of the eigenvectors associated with the Principal Directions (CP) follows. The eigenvalues correspond to the inertia captured by each PC.

## Slide 20
PCA explains the original variables through the strength of their relationship with the factors
The factors are extracted from the data sequentially
The factor that explains the highest % of variance is extracted first
Then, the factor that explains the second highest % of variance is extracted, and so on
The factors are orthogonal to each other
They have zero correlation between them, each representing something unique
Notes: "Eigenvalues are particularly important. The number of factors to retain is a crucial decision in principal component analysis. The goal is to reduce the number of variables to a more manageable number of factors while retaining important information. A good way to make this decision is by examining the eigenvalues. The eigenvalues, one for each factor, add up to the number of variables in the input data. So, if, for example, we have 21 input data variables, the eigenvalues would sum to 21 units of variance. Principal component analysis seeks to explain individual variables through the strength of their relationships with underlying factors. We're looking for factors that explain or underlie more than a single variable. Thus, a good criterion for determining how many factors to retain is if the eigenvalue of a factor is greater than one. If so, it represents more than a single variable, which is why I decided to retain three factors. There are actually six factors in this dataset with eigenvalues greater than one, but the eigenvalues for four, five, and six are only slightly greater than one, so I decided to retain only three factors for the purpose of rotation. There are other ways to make this decision, but all are a mix of objective and subjective criteria.
There are many other reasons why eigenvalues and their accompanying eigenvectors are important. Among these reasons is that the factors are extracted from the data sequentially. The factor that explains the most variance is extracted first. Then, the factor that explains the second-highest degree of variance is extracted, and so on. Furthermore, the factors are all orthogonal to each other. That simply means that the factors have zero correlation with each other, and, in turn, that means that each factor measures something unique. There is no ambiguity about how variance is assigned to each factor. In an orthogonal solution, like the one we are working with here, any variation associated with one factor cannot be associated with any other factor. This is because of how the factors are extracted from the original set of variables.
How to perform PCA 1. Scale and center the original variables by the mean.Calculate the components, scores, and variances.
2. Next, apply the variance threshold so that only the high-variance components are retained. Another common approach is to set a percentage of the total variance to preserve, and then retain the minimum number of components that satisfy this condition
Apply the variance filter. Discard components below the variance threshold. Retain the minimum number of components that cover the target variance.
The scores of the retained components are a new set of features.

## Slide 21: Factor rotation
Ideal
Real
The variables that are weakly related to the component are located near the center of the graph
Notes: Ideally, we would like to have a subset of variables that are strongly related to one of the principal components, and then another subset that is strongly related to the other principal component. This ideal situation rarely arises immediately after the factors have been extracted, but we can get closer to that ideal configuration by rotating the axes, keeping the data points where they are but turning the axes so that each variable is close to the center of the graph on one or more axes and far from the center of the graph only on one axis. Once you have rotated the axes to that alignment, you will find individual variables that are strongly related to a component far from the center of the graph and you will find variables that are weakly related to the component near the origin point of its axes.

## Slide 22: Loadings
Correlation coefficients
-1.0 : Perfect negative correlation.
0.0: No correlation.
1.0: Perfect positive correlation.
In yellow: Loadings>0.5
The variables that load strongly on the same factor may share something in common. For example, products C, D, J, and U load strongly on the first factor. In this example, they are antibiotics
Notes: Only three factors are displayed in columns B, C, and D. Remember that when I completed the principal components dialog box, I wanted to retain three factors. The original variables are shown in column A. The numbers in this table are called factor loadings. You can interpret them as if they were correlation coefficients ranging from a minimum of -1.0, indicating a perfect negative relationship, to 0.0, indicating no relationship, to 1.0, indicating a perfect positive relationship.
I have highlighted in yellow the factor loadings that exceed 0.5. These are stronger loadings, so you can see that products C, D, J, and U load quite strongly on the first factor. Similarly, product H, product I, and product L load quite strongly on the second factor. Four additional products, O, P, R, and S, load moderately on the third factor. Note that there are several products that start with products A, B, E, F, and G, and so on, that do not load strongly on any of the three factors. These products tend to exhibit unique variation, variation that they do not share with other products. You may well find that variables that load strongly on the same factor have something in common. For example, when the components are rotated, products C, D, J, and U load strongly on the first factor. The four products may be, for example, antibiotics. Products H, I, and L, which load strongly on the second factor, could be statins. At this point, if you feel comfortable with the principal component analysis so far, you can start analyzing individual records using factor scores instead of variable scores. In other words, you can act as if the factors were observed variables and analyze different groups or clusters, based on a clustering analysis, according to their factor scores. Products A and B show unique variance, variance they do not share with other products. In reality, it's about understanding what variance these products have in common, what variance they share. That shared variance among products, or samples, is expressed in the factors."

## Slide 23: Outlier detection
La distancia de Hotelling or T2 corresponde a la distancia entre la proyección del individuo sobre el modelo y el centro del modelo
La distancia Q or residual
Notes: "SCORE PLOT: Individual factor map. Proximity and difference between individuals. Loadings: Represent the original variables in the vector space constructed around the PCs. The distance of the projection of an original variable onto the PCs corresponds to a loading or factorial contribution, meaning its weight or its contribution or influence in building PCs. The cosine of the angle formed between an original variable and a PC corresponds to the correlation between both vectors
Q Residual represents the distance of the individual to the principal plane 1-2 of the model
Q- residual indicates how well each sample conforms to the PCA model. It is a measure of the difference, or residual between a sample and its projection into the principal components retained in the model. T2-Hotelling's distance, is a measure of the variation in each sample within the PCA model.
The first multivariate distance used is called the Hotelling's T-squared (T2) distance. Suppose we have a PCA model with A principal components; Hotelling's T2 distance corresponds to the distance between the individual's projection onto the model and the center of the model. If we look at the figure, in the case of a PCA model with 2 PCs, T2 corresponds to the distance between the individual's projection onto the principal plane 1-2 and the center of the PCA model. It is called a strong outlier because it is well modeled by the model, meaning it contributes strongly to the construction of the model. If it is removed from the analysis, the PCA model is significantly affected.
The second multivariate distance is called Q or residual distance. It directly represents the distance of the individual to the principal plane 1-2 of the model. An outlier with a high residual is poorly modeled by a PCA. If it is removed from the PCA model, the impact on the model will be low."

## Slide 24: Part 4: Pretreatments on Raw Spectra

## Slide 25: Near Infrared Spectroscopy
.
Lambert-Beer Law
Ideal conditions
Real conditions
Notes: Ideal conditions for IR spectroscopy correspond to the measurement of a transparent and low-concentration solution.
Under these conditions, for a completely transparent and low-concentration solution, the concentration of a compound of interest is directly proportional to its absorbance. The Lambert-Beer Law states that the absorbance of a compound at a given wavelength will be proportional to the concentration of the component, multiplied by the path length, and the molar absorptivity at that same wavelength, which is characteristic of each compound.
In FT-IR spectroscopy, this law applies to different wavelength values in the infrared spectrum. The different absorbances of the same compound as a function of its wavelengths can be plotted in the figure
"The Beer-Lambert law is never used under ideal conditions. Its application is intended in other raw materials, such as flour, wood, and soil. The relationship between the spectrum and concentration is even more complicated. The Beer-Lambert law is not verified because the measured spectrum is distorted by numerous phenomena.
These phenomena are related to the concentrations of other compounds in the mixture, temperature, sample granularity, etc.
A very visual example is that of a green-colored glass, crushed to different granularities. We observe in the figure that the more crushed the glass is, the whiter it appears.
The color that appears in the equipment is directly proportional to the measured absorbance. An object that appears green has less absorbance in green (525nm) than in blue (480nm) and red (650nm). With a thicker glass, you observe less green and more blue and red.
The absorbance spectra measured for the three colors will be different, even though their chemical content and dye content are identical."

## Slide 26: Why use pretreatments?
The application of a linear model is translated into a dot product of the spectral vector by a vector consisting of the model coefficients
The assignment of a sample to a class is usually done by calculating the distance between the spectral vector and the vectors representing the class centers
All deformations of the spectra entail deformations of the vectors and can:
Distort the space described by the spectra, as seen, for example, in the training set, leading to a poor model estimation
Pretreatments eliminate spectral deformations to approach the pure contribution of the studied parameter
Notes: "The measured spectra are expressed in a column of P variables, meaning vectors of dimension P. Chemometric treatments utilize these vectors in algebraic and numerical calculations
They are used to remove non-chemical deviations from spectral data and prepare the data for subsequent treatments. Pretreatments do NOT alter or affect the original data."
"NIR spectra are influenced by various parameters.
Variation in the chemical and physical properties of samples, as well as the measurement process and changes in the spectrometer, have an influence on the spectra.
Samples measured using diffuse reflectance often exhibit significant differences in the spectra due to the non-uniform distribution of particles.
A practical issue in the application of NIR technique is that spectra are often disrupted by various interferences in a signal acquisition process. The path length of light is often affected by sample thickness, measurement geometry (for non-scattering samples), or physical properties of samples, such as particle/droplet/fiber size and distribution (turbid samples) [4]. Therefore, preprocessing is a critical step in NIR spectral modeling. The goal of preprocessing is to appropriately minimize the unwanted influence of the physical attributes of samples on the measured spectra.“
ALWAYS SAVE THE ORIGINAL SPECTRUM, HAVE A GOOD REASON, UNDERSTAND HOW THE ALGORITH WORKS, RECORD EVERYTHING

## Slide 27: Real spectrum may be influenced by:
Multiplicative and additive effects
Measurement noise
Enlarges the medium length of the optical path by factor k
2) Will cause a certain number of photons to escape from the captor, creating a leak term
It is due to a set of random phenomena present throughout the measurement chain
Efecto aditivo
Efecto aditivo
Addiive effect
Photon diffusion
Notes: En el caso de una muestra real, el espectro estará perturbado por la difusion de fotones y por el ruido de medición.
La difusión de los fotones alarga la longitud media del paso óptico por un factor k (multiplicativo), y genera que cierto número de fotones se escape del captor, esto agrega un término de fuga. El ruido de medicion se debe a un conjunto de fenómenos aleatorios presentes todo a lo largo de la cadena de medición
LUZ PARASITA, RUIDO ELECTRONICO, ETC
Un termino multiplicativo y dos terminus aditivos.
Jean Michel Roger in 2019 explains how a signal is disturbed in the case of a real sample, when its spectrum is damaged by photon scattering and measurement noise:
“In the case of a real sample, the spectrum is disturbed by photon scattering and measurement noise. Photon scattering (1) lengthens the average optical path length by a factor k (multiplicative effect); (2) leads a certain number of photons to escape from the sensor, which adds a leakage term per wavelength l (additive effect): Af (λ). Measurement noise is due to a set of random phenomena present throughout the measurement chain (stray light, electronic noise from the sensor, etc.). It results in an another additive term per wavelength Ab(λ)
Finally, the absorbance actually measured is assigned a multiplicative term and two additive terms:
A(λ) = kε(λ)LC + Af (λ) + Ab(λ) (Equation 1)
Where, A: Absorbance; λ: Wavelength; A(λ): Absorbance at a particular wavelength
k: multiplicative effect; ε: Molar absorption coefficient (M-1 cm-1); L: Optical path length (cm)
C: Molar concentration; Af (λ). Additive effect at a particular wavelength caused by a certain number of photons escaping from the sensor; Ab(λ). Additive effect at a particular wavelength caused by random phenomena”.
The different preprocessing strategies will consist in correcting as many of the three terms k, Af (λ), Ab(λ) as possible. (Jean-Michel Roger M. E., 2019)
It is important to know whether the data has been affected by a multiplicative effect or an additive effect or both. The preprocessing method to be used depends on its efficiency to remove the noise present in NIR spectra. For this reason, it is important to understand firstly very well the type of noise/effect present in each particular case.
It should be mentioned that it is possible to apply more than one pre-treatment sequentially. For example, a multiplicative scaling is often used together with a derivative.

## Slide 28: How to distinguish additive and multiplicative effects
NIR spectra of water and oil mixtures showing a multiplicative effect. Source: WUR.
NIR spectra showing an additive effect.
Source: WUR.
Notes: Figure 1 shows an example of a pure multiplicative effect on mixtures of water and oil spectra. As shown by Jean Michel in his 2019 grapes dataset (Roger, 2019), the multiplicative effect shows the greater the intensity, the greater the distribution of the spectra. Figure 1 shows in the left zone close to 1133.6 nm (x-axis) that the intensity levels have been little affected, whereas the intensity of the spectra in the central zone between 1197 and 1261 nm are more expanded.
When spectra are clearly affected by light scattering, an increase in light intensity with wavelengths in a certain spectral range can be observed. As mentioned above, the light does not fully impact the material to be analyzed, but is lost in the inter-particle spaces. This can happen when the particles do not have the same size, or when there is porosity or gases present in the sample (https://www.aunir.com/light-bites/illuminating-spectra-part-7/). Figure 2 shows an increasing baseline in an spectrum of fish bone, from 5000 cm-1 to 4000 cm-1
Figure 3 shows an example of a pure additive effect in spectra: the dispersion between spectra is the same in the studied spectral range. The effect is strong in areas of weak absorbance (1150 nm), as well as in areas of strong absorbance (1575 nm).
The effects observed in Figures 2 and 3 can be combined, giving rise to an additive effect with increasing baseline of type aX + b, where the dispersion increases with the level of the spectrum and not with the wavelength.
Jean Michel shows spectra with increasing baselines in a Fictional example in Figure 4 (Roger, 2019): “The mean spectrum of the maize dataset was replicated 80 times and each replicate was assigned a baseline of equation aλ + b, with a and b positive but random. This is not a multiplicative effect”.
Analogously, Figure 5 shows one handmade example of NIR spectra with an increasing baseline, this time with multiplicative effect added as well. The mean spectrum was repeated 15 times, each time with increasing effects.

## Slide 29: How to distinguish additive and multiplicative effects
(Fictional) example of an additive effect of
a rising baseline. Source: IRSTEA Montpellier
Handmade example of NIR spectra with combined effect (additive and multiplicative) Source: WUR

## Slide 30: How to distinguish additive and multiplicative effects
A millefeuille shape observed when plotting NIR spectra with additive effect versus the mean spectrum. Source: WUR.
A cone shape observed when plotting
spectra with multiplicative versus their mean
spectrum. Source: from author

## Slide 31: How to distinguish additive and multiplicative effects
No clear cone neither millefeuille shape shown; difficult to tell if there is additive or multiplicative effect present in spectra. Source: WUR
Intensities from the black rectangle can be plotted versus their mean spectrum, to understand the type of effect present. Source: WUR
This shows an additive effect in that spectral range. Source: WUR.

## Slide 32
How to reduce a multiplicative effect (k)?
Logaritmo
Normalization
Multiplicative
Additive
Maximum
Sum
Average
Norm
Standard Deviation
Standard Normal Variate
Notes: La transformación logarítmica> transforma un efecto multipicativo en un efecto aditivo. Se necesita que x sea estrictamente positive. Por ejemplo, el logaritmo, no puede ser aplicado despues de una derivada. |Esta es la transformación natural de Lambert-Beer.
Se trata de divider el espectro, por un parámetro estadístico que sea sensible al efecto multiplicativo en particular.
Se recomineda evitar divider el espectro por un valor único como el punto máximo, ya que puede tener errores.

## Slide 33
To remove noise
Savitsky Golay
w (odd)>d
polynomial of degree d
At each point i in the spectrum, the raw value xi is replaced by zi from a polynomial fitted over a window around point i.
Notes: Ese tipo de filtro puede ponerse en práctica con diferentes algoritmos
La media o mediana móvil: cada punto del espectro se reemplaza por la media o la mediana de los valores del espectro en una ventana alrededor de i.
El ancho de esta ventana es un parámetro a setear. Cuanto más ancha es la ventana, más liso será el espectro.
La media está adaptada, para quitar el ruido de la media nula, como el ruido gausiano. Para quitar los picos aleatorios de media no nula, como los accidentes de numerización, se usa la mediana.
MEDIA
La media aritmética o promedio, se calcula mediante la suma de todos los valores observados y dividiendo por el número de observaciones. La media es la forma más sencilla de resumir una sola variable y por lo general es la mejor medida de tendencia central para los propósitos de la inferencia estadística.
MEDIANA
La mediana es el valor medio o central de un conjunto de observaciones. Para encontrar la mediana, los valores deben estar ordenados. Si el total de elementos del conjunto de datos es par, MEDIANA calcula el promedio de los números centrales. Si el total de elementos es un número impar entonces la MEDIANA tomara el valor ubicado en el centro de la lista. La mediana es útil para variables como los precios de la edad, los ingresos, el volumen de negocios y la vivienda.
MODA
La moda es la incidencia más comúnmente encontrado en un conjunto de datos. Un conjunto de datos puede tener más de una moda. La moda es útil cuando el material más común, característica o valor de un conjunto de datos se requiere.
Se trata de un alisado polynomial aplicado sobre una ventana móvil
En cada punto I del espectro, el valor bruto xi se reemplaza por zi de un polinomio ajustado sobre una ventana alrededor del punto i.
ATENCION> EL espectro completo no se reemplaza por el polinomio. ES LA ORDENADA CENTRAL DE LA VENTANA QUE SE REEMPLAZA POR LA ORDENADA CENTRAL DE POLINOMIO.
Dos parámetros deben ser seteados para este algoritmo: el ancho de la ventana (w) y el grado del polinomio (d). Cuanto más grande sea w, más liso será el espectro. Cuanto mayor sea el grado del polinomio, menos liso será el espectro. Por razones de cálculo, w debe ser impar y w>d
El suavizado de espectros
Aumenta la relación Señal Ruido (S/N)
Si esta demasiado suavizado el espectro se deforma y pierde información.
Algoritmo es una convolución entre el espectro y un vector que define el grado de suavizado

## Slide 34
To remove noise
Detrend
Delete the global tendency from the spectrum, modeled by a polynomium
Order 0
Average
Orden 1
Line
Orden 2
Parabola
The residuals that remain after removing the line correspond to the absorbance peaks related to the chemical components of the sample
Spectrum of Wheat. Source: IRSTEA Montpellier
Notes: Efecto aditivo debido a la fuga por diffusion de protones.
LINEA DE BASE CRECIENTE DEBIDO A LA DIFUSION DE LA LUZ
Para reducir la lineas de base, se usa detrend (tendencia).
Tendencia creciente con el número de longitudes de onda. Este fenómeno no se debe a la química, pero a la diffusion de la luz, cuya intensidad crece con la longitude de onda (en este rango de longitudes de onda). Este espectro de compone de una línea de base, sobre la cual se superpone un espectro de absorbancia.
.
DETREND- Borrar del espectro la tendencia global, modelizada por un polinomio.
Detrend de orden cero: El polinomio que se elimina es una constant que corresponde al promedio del espectro.
Orden 1> Se identifica por regression lineal, la recta que se ajusta lo mejor possible a los espectros, y se elimina. En el caso de este espectro de maíz, los residuos que quedan de esta recta, son los picos de absorcion que corresponden a los componentes quimicos de la muestra.
Orden 2- Se identifica una parabola, o una función cúbica, se la elimina del espectro.

## Slide 35
Detrend
Notes: La figura muestra el efecto de la correccion detrend en los espectros. Se ve claramente que el orden cer reduce las translaciones verticals puras. El orden 1 reduce las lineas de base más finas, como esa ecuación (recta). El orden dos no aporta gran cosa.

## Slide 36: Derivatives
They are also used to decrease the baseline. If the baselines are polynomials of degree K, then the derivatives of order k+1 will make them disappear
Notes: El concepto de derivada de una función matemática se halla íntimamente relacionado con la noción de límite. Así, la derivada se entiende como la variación que experimenta la función de forma instantánea, es decir, entre cada dos puntos de su dominio suficientemente próximos entre sí.
Tb se utiliza para DISMINUIR LAS LINEAS DE BASE. Ubiquémosnos en el punto de vista de tratamiento de señal, donde el espectro se toma como una función de la longitud de onda.
En NIR, las líneas de base son lineales, a y b sn coeficientes incognitas y especificos para cada muestra. En esta nueva expression derivada, se encuentra la proporcionalidad con C, lo cual es necesario para hacer una calibracón. Sn embargo, el espectro resultante de esta transformación, no tiene nada que ver con una absorbancia. La derivada solo elimina la linea de base.
Como e está constituida por picos de absorcion, una derivada primera va a transformer estos picos en cero.
Savitzky-Golay algorithm (SG) is widely used for spectral derivative of near-infrared spectra [9]. Local least squares regression is performed in SG derivation, providing high computational efficiency of the algorithm.
A first-order derivative can eliminate constant baseline (offset) and a second-order derivative can also eliminate slope.
Los puntos máximos y mínimos de los picos originales tienen un valor Y de cero en la primera derivada.
Los puntos máximos y mínimos en la primera curva derivada son los puntos de la tasa máxima de cambio
La primera derivada es útil para revelar picos que aparecen como hombros en los espectros originales.
Muestra la tasa de cambio en todo el espectro.
Los hombros se vuelven más estrechos y, por lo tanto, son más fáciles de ver.
Los puntos máximos y mínimos en la primera curva derivada son los puntos de la tasa máxima de cambio y no los puntos máximos y mínimos de los picos originales tienen un valor Y de cero en la primera derivada.
Como e está constituida por picos de absorcion, una derivada primera va a transformer estos picos en cero (línea punteada de la figura)
Y va a amplificar las zonas de pendiente máxima, (trazo mixto en la figura).
SE AMPLIFICAN LAS ZONAS DE PENDIENTE MAXIMA
LOS PICOS DE ABSORCION SE TRANSFORMAN EN CERO. Puntos de inflexion van a cero.
Use la segunda derivada para encontrar la ubicación exacta (centro) de los hombros en los espectros originales.
Use la segunda derivada para encontrar la ubicación exacta (centro) de los hombros en los espectros originales.
Use la segunda derivada para encontrar la ubicación exacta (centro) de los hombros en los espectros originales.
Bandas significativamente mas estrechas
Los picos aparecen en las mismas ubicaciones que los picos del espectro original
La segunda derivada tiene más ruido de referencia que la primera derivada. Para cada operación derivada que realice, el nivel de ruido aumenta ligeramente, la intensidad de la señal disminuye drásticamente y la relación señal / ruido disminuye.
Por el contrario, una derivada segunda, tendrá tendencia a magnificar los picos. Supongamos que medimos un espectro de una muestra que está compuesta por una mezcla de dos components cuyos espectros individuales conocidos. Es dificil cuantificar los picos de cada componente, debido a su recubrimiento. La aplicacion de la derivada segunda permite la deconvolución de los 2 picos de la mezcla.
Sin embargo, la derivación numérica de una señal es una operación delicada, ya que ella disminuye considerablemente la relación señal ruido.
En efecto, cuando se calcula la derivada, se hace la sustracción de los valores vecinos *en (i y e i+1). Si en estos valores existen ruidos aleatorios, y por lo tanto independientes, la varianza resultante es la suma de las varianzas de xi y xi+1. El ruido aumenta en raíz de dos.
Es por esto que la derivación de espectros NIRse hace con algoritmos que alisan y derivan al mismo tiempo
In fact, when the derivative is calculated, subtraction of neighboring values * is made in (i and e i + 1). If there are random, and therefore independent, noises in these values, the resulting variance is the sum of the variances of xi and xi + 1. The noise increases in root of two.
This is why the derivation of NIR spectra is done with algorithms that smooth and derive at the same time

## Slide 37: Standard Normal Variate (SNV)
NIR spectra of honey, with both additive and multiplicative effects. Source: SENASA (Argentina)
SNV NIR spectra of honey. Additive and multiplicative effects are gone. Source: SENASA (Argentina)
The SNV pre-processing of a spectrum X consists of removing the mean of X at each of its points, then dividing them by the standard deviation of X.
Notes: The Standard Normal Variate algorithm (SNV) allows to process both additive effect and multiplicative effects (R.J.Barnes, 1989). (R.J.Barnes, Correction to the description of standard normal variat (snv) and detrend (dt) transformations in practical spectroscopy with applications in food and beverage analysis,, 1993)
Jean Michel Roger explains it in his report in written in 2019:
“The SNV pre-processing of a spectrum X consists of removing the mean of X at each of its points, then dividing them by the standard deviation of X. This is a centering/ reduction calculation, but performed on the rows of the spectra matrix, instead of the columns. SNV is the combination of a Detrend of order 0 and a normalization by the standard deviation. As a consequence, SNV standardises each spectrum to have mean 0 and standard deviation 1 (Roger, 2019).
This preprocessing can be applied to a single spectrum and does not require any parameter setting. The SNV algorithm is given by the following sequence:
Xc = X – mean(X)
XSNV = Xc / SD(Xc)
Where 1 is a P-dimensional vector containing only 1s and P is the number of wavelengths.
Note: If you apply 1st derivative in first, it leads to a constant spectrum and then SNV is even not applicable ; but you can apply SNV first and then derivative”. (Roger, 2019)

## Slide 38: Multiplicative Scatter Correction (MSC)
Mathematically, if we call Xm​ the mean spectrum, the multiplicative scatter correction is done in two steps.
We first regress each spectrum Xi​ against the mean spectrum. This is done by ordinary least squares: Xi​≈ai​+bi​Xm​ .
We calculate the corrected spectrum msc=()/Ximsc​=(Xi​−ai​)/bi​.
Notes: Multiplicative Scatter Correction (Geladi et al., 1985), as its name indicates, is typically used to compensate for light scattering effects and changes in path length. Indeed, light scattering from solid samples and emulsions can result in multiplicative deviations that are dependent on the wavelength (Heise & Winzen, 1985). MSC minimizes these deviations by fitting a linear model between a reference spectrum and other spectra of the dataset using the linear least squares method. The reference spectrum is often chosen as the average of all spectra in the dataset. The model coefficients are used to compute the ideal (i.e. msc corrected) spectra. Following application of MSC, all spctra appear to have the same absorbance level. MSC and SNV usually lead to similar results. However, as mentioned, MSC requires the use of a reference spectrum while SNV does not.
MSC requires a reference spectrum to start with. This is the most important difference between MSC and SNV. The reference spectrum is ideally a spectrum free of scattering effects.
Now, as you can gather, getting our hands on a spectrum that is free of unwanted scattering effects is not easy, definitely not across all wavelengths we are interested in. For this reason, if the data is reasonably well behaved, we can take the average spectrum to be a close approximation to the ideal spectrum we are after. Particle size and path length effects should vary randomly from sample to sample, and therefore the average should reasonably reduce these effects, at least in the approximations that these effects are genuinely random. This is the main assumption behind MSC.

## Slide 39: Extended Multiplicative Signal Correction (EMSC)
Notes: MSC: an approximation of each spectrum by its average spectrum and a constant baseline. Sometimes this is not sufficient.
. It is interesting to note that in Eq. (8) the interesting part, namely the biochemical differences among the spectra, is expressed by the residual
The part of the spectra that is not taken into account by the model is assigned to the term eð Þ ν~ in Eq. (10). Thus, the residual eð Þ ν~ contains the interesting chemical differences between the samples.
e, in absorption spectroscopy the multiplicative effect is proportional to the effective optical path length
The framework of EMSC allows, apart from basic preprocessing as baseline correction and normalization, to separate and quantify different types of chemical and physical variations in vibrational spectra. EMSC thus enables the user to study these different effects separately.
Preprocessing methods in vibrational spectroscopy may be divided into two categories, namely filtering methods and model-based methods. The filtering methods, like calculating second derivatives or performing vector normalization, simply transform spectra into a presumably “better” version of the same data, removing some undesired types of variation. Model-based methods, on the other hand, allow for quantifying and separating between the different types of physical and chemical variations in the spectra. Hence, the information filtered is not lost, but statistical estimates of the mathematical parameters involved in the filtering are obtained and can be used to study the different effects separately
EMSC has been shown to be a reliable tool to correct for additive baseline effects, multiplicative scaling effects, and interference effects.

## Slide 41: Part 5: Feature selection methods
"Aristotle: Nature operates in the shortest way possible"

## Slide 42: Feature selection
Increasing the number of variables:
Introduces unnecessary NOISE for discrimination, especially if they are strongly correlated,
Carries a risk of OVERFITTING the models.
Using a simple model with few variables has a better chance of being generalized to a new sample than a model with hundreds of variables, which may fit the training set perfectly but has limited generalization power.
3 Approaches:
Variable Transformation and/or Selection
Discretization
Group Summary
Notes: Además, si se construye un modelo con la intensión de ser utilizado en rutina, es generalmente más eficaz y económico, tener que medir solo algunas variables.
A pesar de que algunos métodos como PLS puede tolerar la redundancia de variables.
Predictive models use relationships between one or more features to predict the value of a response feature, such as power.
However, including additional predictors like the color of the car that are unrelated to the response will not provide additional benefits and may even affect the model's performance.
The goal of feature selection is to enhance your model by retaining only the features that add value and discarding the rest.
The resulting model will also be simpler to interpret and more computationally efficient.

## Slide 43: Variable transformation
TRANSFORMACION DE VARIABLES
This involves applying an equation to existing variables to create a new feature
No clear correlation
Example: Is it possible to accurately classify the health status of each individual from the original variables of age, height, weight, and location?
BMI= Weight (Kg)
Height (m)2
Correlation
Notes: For example, consider this dataset containing information about hospital patients, with one row for each individual.
The initial question is, "Is it possible to accurately classify the health status of each individual from the raw variables of age, height, weight, and location?" To start, we can input the raw variables into a model and attempt to compare, for instance, health status with weight or height.
However, a quick visualization reveals that these variables on their own are not good for predicting health status because there is no clear correlation. So, what can we do?
Well, you can begin the process of feature engineering. It is possible to enhance a model's predictive power by applying domain knowledge to generate useful features. There might be many features that are better predictors of a patient's health status than the original variables on their own. Visualizing the BMI (Body Mass Index) shows that it is indeed more correlated with the patient's health status than weight or height alone.
The second approach of discretization involves dividing the data into groups using a numerical variable. This can be helpful because sometimes the important aspect of a variable is not its exact value but the range in which it exists.
For instance, the age variable naturally lends itself to grouping patients into three categories: child, adult, and senior. You can then add this categorical variable as a new feature to the table.
Discretization is the process by which we can transform continuous variables, models, or functions into a discrete form. This is done by creating a set of continuous intervals (bins) that span the range of our desired variable, model, or function.
Why Discretization is Important: Mathematical problems with continuous data have an infinite number of degrees of freedom. Sometimes, it's necessary to limit these degrees of freedom because calculations cannot go on indefinitely.
It is often easier to understand continuous data (like weight) when they are divided and stored in meaningful categories or groups.
Finally, the third approach is to summarize groups within your data. For example, how does a patient's weight compare to the rest of their age group? To find out, calculate the average weight for each category.
Then, append this result to your table as a new variable. Next, you can calculate the difference between each individual's weight and the average.
Each of these approaches can be applied to any dataset. However, the exact implementation depends on your knowledge of the data itself. In theory, you can generate an infinite number of new features, but only a few will prove to be truly useful.

## Slide 44
Variable selection methods
Filter
Wrapping
Embedded
Notes: Hay tres tipos de métodos de selección de funciones, envoltorio, incrustado y filtro.
FILTER METHODS.
Filter methods are used before training the model. They are based on the statistical properties of the predictors and their relationship with the response.
FILTERS: Filters are used for a preliminary variable selection before more sophisticated methods. That is, they are typically used in combination with wrapper and embedded methods.
They involve evaluating the intrinsic characteristics of the data independently of the specific problem considered. They can be applied regardless of the model.
Select a significance level α (e.g., α = 0.20 or α = 0.157). • Conduct K univariate models. • Use all variables in a multivariate model with a univariate p-value < α. • Sometimes followed by backward elimination.
Simple method, can be calculated by any computer, but it only works properly with UNCORRELATED VARIABLES
A Filter method Example: Variance Thresholds
Here, we simply compute the variance of each feature, and we select the subset of features based on a user-specified threshold. E.g., “keep all features that have a variance greater or equal to x” or “keep the the top k features with the largest variance.” We assume that features with a higher variance may contain more useful information, but note that we are not taking the relationship between feature variables or feature and target variables into account, which is one of the drawbacks of filter methods.
Forward Selection:
Select a significance level α1.
Estimate the model.
Repeat:
  While the most significant term has p < α, add it and re-estimate. Variant: Forward, in steps.
Select two significance levels α1 and α2.
Repeat:
  While the most significant term excluded has p < α1, add it and re-estimate.
  While the least significant term included has p > α2, remove it and re-estimate.
Backward Selection:
Select a significance level α2.
Estimate the full model.
Repeat:
  While the least significant term has p > α2, remove it and re-estimate. Variant: Backward, in steps.
Select two significance levels α1 and α2.
Repeat:
  While the least significant term included has p > α2, remove it and re-estimate.
  If the most significant term excluded has p < α1, add it and re-estimate.

## Slide 45
These variable selection methods depend on model performance. It is an iterative process of back and forth, where the chosen features are evaluated in relation to the final model performance in sequential stages.
Train the model
Evaluation of model performance
Variable selection
Variable
Variable
Variable
Best features
Optimized model
Wrapping
methods
Notes: Wrapper methods train a new model for each subset and use the model's error rate on a hold-out set to score feature subsets.
When wrapper methods are used, variable selection is performed simultaneously with learning. This involves an iterative process between feature subset selection stages and subset evaluation stages using the chosen learning method.
The forward method works as follows:
Without loss of generality, let's consider a regression problem. It starts by selecting the variable that best addresses the problem, meaning the variable most correlated with the target variable. Then, it looks for the variable that is most associated with the first one through a multiple regression. This process continues until all variables are introduced. Different criteria are used to choose the best model among the stages performed.
To prevent overfitting: Cross-validation.
we have to train and cross-validate our model for each feature subset combination

## Slide 46: Covariance
Indicates the degree of joint variation between two random variables with respect to their means
It can be used to understand the direction of the relationship between two variables
The correlation coefficient is equal to the covariance divided by the product of the standard deviations of the variables
Notes: La covarianza mide la relación lineal entre dos variables. Aunque la covarianza es similar a la correlación entre dos variables, difieren de las siguientes maneras:Los coeficientes de correlación están estandarizados. Por lo tanto, una relación lineal perfecta da como resultado un coeficiente de 1. La correlación mide tanto la fuerza como la dirección de la relación lineal entre dos variables.
Los valores de covarianza no están estandarizados. Por consiguiente, la covarianza puede ir desde infinito negativo hasta infinito positivo. Por lo tanto, el valor de una relación lineal perfecta depende de los datos. Puesto que los datos no están estandarizados, es difícil determinar la fuerza de la relación entre las variables.
Los valores de covarianza positivos indican que los valores por encima del promedio de una variable están asociados con los valores por encima del promedio de la otra variable y los valores por debajo del promedio están asociados de manera similar. Los valores de covarianza negativos indican que los valores por encima del promedio de una variable están asociados con los valores por debajo del promedio de la otra variable.
Por lo tanto, una covarianza positiva siempre producirá una correlación positiva y una covarianza negativa siempre generará una correlación negativa

## Slide 47: Wrapping method: Covariance selection (CovSel)
In each iteration, one variable X is selected on a criterion of the maximization of the covariances with Y
Once the variable with the highest covariance is isolated and selected, all other predictive factors and responses are orthogonalized with respect to it, and the process is repeated until the fixed number of variables has been selected.
Notes: CovSel is a variable selection method designed to identify the most relevant features in a regression context. As the name suggests, the relevance of each variable is assessed by estimating the covariance between each predictor and the responses. Once the variable with the highest covariance is singled out and selected, all other predictive factors and responses are orthogonalized with respect to it, and the procedure is repeated until a fixed number of variables has been selected
CovSel aims to iteratively select variables from a predictive variable matrix X with N rows and P columns, which have good predictive capability for the variables combined in a matrix Y with N rows and Q columns. The variables selected in the end will be relatively few in relation to P. The most commonly encountered case is when a single variable needs to be predicted (Q=1), but here we will address the general case where Q can potentially be greater than 1. In each iteration, the variable X is selected based on the criterion of maximizing covariances with Y

## Slide 48: Covariance selection (CovSel)
The selected variables from X should have:
Good predictive power for Y
The highest possible variability
Notes: El criterio de la maximización de las covarianzas reposa sobre la idea de que las variables de X seleccionadas deben presenter dos características:
Deben presenter un buen poder predictivo de Y
Deben presenter la mayor variabiliad possible. De la misma manera que en la regression PLS, se tiene interés en dar la mayor importancia a las variables en X que son las más intensas. Como el ruido aleatorio depende en general de sus intensidades, la elección de una variable intensa aporta menos importancia relativa al ruido.

## Slide 49: Covariance selection (CovSel)
1) The number of variables (L) from matrix X is chosen
2) X is centered and Y is centered and reduced
3) Covariance is calculated between each variable Xp and Yp
4) The predictor with highest covariance wth Yp is chosen
5) Xp is orthogonalized with respect to the chosen response
6) Yp is orthogonalized with respect to the chosen preditor
Notes: The first stage of CovSel involves evaluating the covariance between the columns of X and Y.
The output matrix X is centered, and the matrix Y is centered and scaled (the intensities of the Y variables are not intended to play a role in the selection).
All covariances between the columns of X and Y can be directly calculated as follows:
The covariance matrix C has dimensions QxP. An element Cij of C gives the covariance between the variable with index I of Y and the variable with index j of X.
The number of variables (L) to be selected from the Xdata block is defined.
The Xblock is centered on the mean, and Y is autoscaled (resulting in Xp and Yp, respectively).
The covariance between each variable Xp (xp, i) and Yp is calculated.
The predictor xSel, l with the highest covariance with Yp is selected.
Xp is orthogonalized with respect to the selected variable xSel, resulting in XO.
After the sixth step, you start again from the third step, replacing Xp with XO, and repeating all the steps L-1 times.
Note that in step 4, Xsel is selected but not removed from Xp.
Also, note that the CovSel algorithm follows the same pattern as the NIPALS PLS. In both, covariance is maximized, but in CovSel, the construction of a latent variable is replaced by variable selection.

## Slide 50
Train the model
Feature selection
Variable
Variable
Variable
Useful variables
Optimized model
Useless features
Embedded
methods
They automatically perform feature selection as part of the model training
The result is a trained model that highlights the useful features and disregards the rest
Notes: Integrated methods, or a class of machine learning models, automatically perform feature selection as part of the model training process.
The result is a trained model that emphasizes useful features and disregards the rest.
Both wrapper methods and integrated methods rely on model performance to guide feature selection.
Wrapper methods consider unimportant features iteratively based on the evaluation metric, while Embedded methods perform feature selection and training of the algorithm in parallel. In other words, the feature selection process is an integral part of the classification/regressor model.
An Embedded Method Example: L1 Regularization
L1 (or LASSO) regression for generalized linear models can be understood as adding a penalty against complexity to reduce the degree of overfitting or variance of a model by adding more bias. Here, we add a penalty term directly to the cost function,
regularized_cost = cost + regularization_penalty
In L1 regularization, the penalty term is
L1 : λ Σkiwi= λw1,where w is our k-dimensional feature vector. Through adding the L1 term, our objective function now becomes the minimization of the regularized cost, and since the penalty term grows with the value of the weight parameters (λ is just a free parameter to fine-tune the regularization strength), we can induce sparsity through this L1 vector norm, which can be considered as an intrinsic way of feature selection that is part of the model training step.

## Slide 51: Part 6: Cross validation and Data Split

## Slide 52: Some methods for the selection of Representative Learning and Test Sets
Randomly
Kernnand Stone
Onion
Duplex
Reducennsamples
Source:https://eigenvector.com/wpcontent/uploads/2022/10/Onion_SampleSelection.pdf
Kennard-Stone
Onion
Notes: Kernnard-Stone selects a subset of samples which uniformly cover the data set and includes exterior samples as the calibration set. The remainder are placed in the test set.
This method selects samples that best span the same range as the original data, but with an even distribution of samples accross that range.
The Kennard-Stone (KS) algorithm starts by finding two samples in that have the largest Euclidean distance; these are the two samples furthest
apart. Subsequent points are selected that maximize the Euclidean distance from previously selected points. The result is that, for a given fraction (fraction) of the data to be parsed into the calibration set,
samples are selected to smoothly fill the data space. The remaining samples are placed into the test set.
2) Reducennsamples
Select a subset of samples by removing nearest neighbors. . Performs a selection of samples which fill out the multivariate space by removing samples which are similar to each other based on nearest neighbor distance. This algorithm is useful in selectin the minimun number of samples needed to define a subspace and reduce the number of reference measurements needed, or amount of data needed to be stored.
Initially, the nearest neighbor of each sample is found along with the distance between the neighbors. Of the two nearest samples, one is excluded from the data and the distances are recalculated. This process is repeated until the either the smallest distance between samples reaches a maximun limit, or the number of samples reaches a lower limit.
3) Onion selects an exterior set of samples for calibration and the next exterior layer as test samples. The procedure is repeated for non-onion layers with the remainder interior samples randomly assigned to calibration and test sets. This method first selects a ring of the most unique samples (based on distance from the mean), like the outer most layer of an onion. These are used in the calibration set. Next, a ring of less unique samples, just inside the first set (the next onion layer), is put into the validation set. This is repeated two more times, so there are three outer rings of most unique and less unique samples. Finally, all the remaining samples are split randomnly into calibration and validation.
Distance measure: Euclidean does simple, non-scaled distance. Mahalanobis scales each direction by the covariance matrix (correcting for unusually small or large directions).

## Slide 53: Some methods for the selection of Representative Learning and Test Sets
Randomly
Kernnand Stone
Onion
Duplex
Reducennsamples
Source:https://eigenvector.com/wpcontent/uploads/2022/10/Onion_SampleSelection.pdf
Duplex
Randomly
Notes: Kernnard-Stone selects a subset of samples which uniformly cover the data set and includes exterior samples as the calibration set. The remainder are placed in the test set.
This method selects samples that best span the same range as the original data, but with an even distribution of samples accross that range.
The Kennard-Stone (KS) algorithm starts by finding two samples in that have the largest Euclidean distance; these are the two samples furthest
apart. Subsequent points are selected that maximize the Euclidean distance from previously selected points. The result is that, for a given fraction (fraction) of the data to be parsed into the calibration set,
samples are selected to smoothly fill the data space. The remaining samples are placed into the test set.
2) Reducennsamples
Select a subset of samples by removing nearest neighbors. . Performs a selection of samples which fill out the multivariate space by removing samples which are similar to each other based on nearest neighbor distance. This algorithm is useful in selectin the minimun number of samples needed to define a subspace and reduce the number of reference measurements needed, or amount of data needed to be stored.
Initially, the nearest neighbor of each sample is found along with the distance between the neighbors. Of the two nearest samples, one is excluded from the data and the distances are recalculated. This process is repeated until the either the smallest distance between samples reaches a maximun limit, or the number of samples reaches a lower limit.
3) Onion selects an exterior set of samples for calibration and the next exterior layer as test samples. The procedure is repeated for non-onion layers with the remainder interior samples randomly assigned to calibration and test sets. This method first selects a ring of the most unique samples (based on distance from the mean), like the outer most layer of an onion. These are used in the calibration set. Next, a ring of less unique samples, just inside the first set (the next onion layer), is put into the validation set. This is repeated two more times, so there are three outer rings of most unique and less unique samples. Finally, all the remaining samples are split randomnly into calibration and validation.
Distance measure: Euclidean does simple, non-scaled distance. Mahalanobis scales each direction by the covariance matrix (correcting for unusually small or large directions).

## Slide 54: Cross validation
Leave One Out:
Over-estimates the predictive capacity of the model
Use only when data set has few samples
Notes: "The most commonly used method for determining the optimal number of latent variables is cross-validation, which calculates the accuracy of predictions for models with an increasing number of latent variables. In its simplest implementation, the data matrix is repeatedly split into two parts, one part (calibration samples) is used to build calibration models with different numbers of latent variables, and the other part (validation samples) to calculate prediction errors.
After applying this procedure to all samples, for each of the models with different numbers of latent variables, an average value for the sum of squared prediction errors, or PRESS, is calculated. The optimal dimensionality for the PLS regression model is the one that minimizes this PRESS or RMSECV value.“
"To increase the precision of the predictive model, it is useful to perform the validation experiment multiple times, reassigning observations to different sets.
General outline of cross-validation: In practice, the calibration set is divided into separate segments of the same size. Each segment forms a validation set each time. The division can be done in various ways. One particular case is Leave One Out, where one sample is removed at a time. Leave One Out: Overestimates the predictive capacity of the model."

## Slide 55: Part 7: Calibration and validation of classification models

## Slide 56: Principles and Objectives of Discrimination
We have a data matrix (nxp) where n samples
were measured for p quantitative variables, and a vector Y of size n measured on the same samples
This vector represents the membership of each sample to each class K
Each class contains at least one sample, and each sample belongs to a single class
Notes: "The principal components are NOT related to the response Y.
They are ordered based on the variability of the X variables, but NOT the response Y.
PCA is only an exploratory tool and only represents the variability in X

## Slide 57: Partial Least Squares Discriminant Analysis
Objective: To achieve a linear transformation that maps the data into a lower-dimensional space with the least possible error
Supervised version of PCA
In PCA, the transformation preserves (in its first principal component) the maximum possible variation in the original data
In PLS-DA, the transformation preserves (in its first principal component) the maximum possible covariance between the original data and their labeling
Both can be described as iterative processes in which the error term is used to define the next principal component
Notes: Partial Least Squares Discriminant Analysis (PLS-DA) is a linear classification method that combines the properties of Partial Least Squares regression with the discriminative power of a classification technique. PLS-DA is based on the PLS regression algorithm (PLS1 when dealing with a single dependent variable Y and PLS2 in the presence of multiple dependent variables Y) which seeks latent variables with maximum covariance with the Y variables. Relevant sources of data variability are modeled by these latent variables (LV), which are linear combinations of the original variables, allowing for graphical visualization and understanding of different data patterns and relationships through the scores and loadings of LV.
PLS1 – Variance is extracted individually for each y. A different set of factors for each y.
PLS2 – Variance is related to several y variables simultaneously. Only one set of factors.
PLS-DA can be considered a 'supervised' version of Principal Component Analysis (PCA) in the sense that it achieves dimensionality reduction but with full awareness of class labels. Besides its use in dimensionality reduction, it can be adapted for feature selection as well as for classification. The goal of dimensionality reduction methods like PCA and PLS-DA is to find a linear transformation that maps the data into a lower-dimensional space with minimal error. If we think of the original data matrix as a collection of m-dimensional vectors (i.e., X is an x by m matrix), the above objective can be seen as finding the transformation matrix A of am × d that optimally transforms the data matrix X into a collection of n d-dimensional vectors S. Hence, S = XA + E, where E is the error matrix. The S matrix, with its rows corresponding to the transformed vectors, provides d-dimensional scores for each of the n vectors in X. In PCA, the transformation preserves the maximum possible variation in the original data (in its first principal component). In PLS-DA, the transformation preserves the maximum possible covariance between the original data and its labeling (in its first principal component). Both can be described as iterative processes where the error term is used to define the next principal component. As mentioned earlier, PCA and PLS-DA are seen as unsupervised and supervised counterparts of dimensionality reduction tools. The new features representing the reduced dimensions are known as principal components.
PLS Discriminant Analysis (PLS-DA) is performed to enhance the separation between groups by rotating the PCA (Principal Component Analysis) components to achieve maximum class separation and to understand which variables carry class separation information. PLS-DA consists of a classical PLS regression where the response variable is a category expressing unit class membership. Therefore, PLS-DA does not allow for response variables other than defining the groups of individuals. As a result, all measured variables play the same role in terms of class assignment. In reality, PLS components are created by trying to strike a suitable compromise between two purposes: describing the set of explanatory variables and predicting the response variables. A classification based on PLS should benefit from this property towards constructing typologies with intrinsic predictive power."

## Slide 58: Partial Least Squares Discriminant Analysis
It consists of a classical PLS regression where the response variable is a category expressing the membership of samples in classes
The relevant sources of data variability are modeled by the Latent Variables (LVs) which are linear combinations of the original variables
A fictitious matrix (Y) that records membership with 1s and 0s is combined with a spectral set (X), and PLS is implemented in the normal manner
Notes: It consists of a classical PLS regression where the response variable is a category expressing the membership of samples in classes. All measured variables play the same role with regard to class assignment.“
Rotating the A (Principal Component Analysis) components is also performed in order to achieve maximum separation between classes and to understand which variables carry the class separation information
The PLSDA calculation procedure consists of the following stages:
Construction of the complete disruptive code for the matrix Y associated with Y,
Using the PLS2 method, with the quantitative variables X as predictors and Y as the matrix to predict
DISRUPTIVE CODE: It is observed that to create this complete disruptive code Y associated with y, the matrix has as many columns as classes, and each row contains only one membership. The rank of Y is equal to K-1

## Slide 59: Partial Least Squares Discriminant Analysis
PLS-DA provides estimated values for each sample and for each class. These values will not be exactly 1 or 0; however, if the calculated y is closer to 0, then the sample likely does not belong to that class, while a value closer to 1 would indicate the opposite
To make a class assignment, a threshold can be defined for each class
Thresholds can be calculated on the basis of the Bayes theorem
Notes: PLS-DA devuelve valores estimados para cada muestra y para cada clase. Estos valores no serán exactamente 1 ni 0; sin embargo, si ycalculado está más cerca de 0, entonces la muestra probablemente no pertenezca a esa clase, mientras que un valor más cercano a 1 indicaría lo contrario.
To make a class assignment, a threshold can be defined for each class: if ycalcig is greater than the threshold defined for the g-th class, then the i-th sample is assigned to the g-th class, otherwise not. The threshold is selected at the point where the number of false positives and false negatives is minimized. It can happen that the estimated class values for one sample are higher or lower than all the defined class thresholds. Thus, if higher, the sample would be assigned to all classes and, accordingly, would be a confused sample. Conversely, if lower, the sample would not be recognised as member of any class. In both cases, the sample can be defined as ‘not assigned’.
Thresholds can be calculated on the basis of the Bayes theorem. This assumes that the estimated values follow a relatively normal distribution, which should be comparable to what will be observed for future samples.

## Slide 60
Notes: Latent variables are determined by CROSS VALIDATION

## Slide 61: Part 8: Hands on data

## Slide 62: Hands on data: R code available